## metadata
PAR ID | DOI | Title | Creator/Author | Corporate Author | Editor | Publication Date | Journal Name | Journal Volume | Journal Issue | Page Range or eLocation | ISSN | ISBN | Description/Abstract | Sponsoring Org. | Award IDs | Subject | Publication Year | Size | Format | Version | Resource Type | Publisher | Location | Rights | Institution
10109409 | https://doi.org/10.5194/amt-11-6419-2018 | Can ozone be used to calibrate aerosol photoacoustic spectrometers? | Fischer, D. Al; Smith, Geoffrey D. |  |  | 2018-01-01T05:00:00Z | Atmospheric Measurement Techniques | 11 | 12 | 6419 to 6427 | 1867-8548 |  | <p><strong>Abstract.</strong> Photoacoustic spectroscopy (PAS) has become a popular technique for measuringabsorption of light by atmospheric aerosols in both the laboratory andfield campaigns. It has low detection limits, measures suspended aerosols,and is insensitive to scattering. But PAS requires rigorous calibration to beapplied quantitatively. Often, a PAS instrument is either filled with a gasof known concentration and absorption cross section, such that the absorptionin the cell can be calculated from the product of the two, or the absorptionis measured independently with a technique such as cavity ring-downspectroscopy. Then, the PAS signal can be regressed upon the known absorptionto determine a calibration slope that reflects the sensitivity constant ofthe cell and microphone. Ozone has been used for calibrating PAS instrumentsdue to its well-known UV–visible absorption spectrum and the ease with whichit can be generated. However, it is known to photodissociate up toapproximately 1120<span class="thinspace"></span>nm via the <span class="inline-formula">O<sub>3</sub></span> <span class="inline-formula">+</span> <span class="inline-formula"><math xmlns="http://www.w3.org/1998/Math/MathML" id="M3" display="inline" overflow="scroll" dspmath="mathml"><mrow><mi>h</mi><mi mathvariant="italic">ν</mi><mspace linebreak="nobreak" width="0.25em"/><mo>(</mo><mo>&amp;gt;</mo><mn mathvariant="normal">1.1</mn><mi mathvariant="normal">eV</mi><mo>)</mo><mo>→</mo><mrow class="chem"><msub><mi mathvariant="normal">O</mi><mn mathvariant="normal">2</mn></msub></mrow><msup><mo>(</mo><mn mathvariant="normal">3</mn></msup><msubsup><mi mathvariant="normal">Σ</mi><mi>g</mi><mo>-</mo></msubsup><mo>)</mo></mrow></math><span><svg:svg xmlns:svg="http://www.w3.org/2000/svg" width="117pt" height="17pt" class="svg-formula" dspmath="mathimg" md5hash="2540a17128385cdeb672d3ea6820a4f8"><svg:image xmlns:xlink="http://www.w3.org/1999/xlink" xlink:href="amt-11-6419-2018-ie00001.svg" width="117pt" height="17pt" src="amt-11-6419-2018-ie00001.png"/></svg:svg></span></span> <span class="inline-formula">+</span> <span class="inline-formula">O(<sup>3</sup>P)</span> pathway, which is likely tolead to inaccuracies in aerosol measurements. Two recent studies haveinvestigated the use of <span class="inline-formula">O<sub>3</sub></span> for PAS calibration but have reachedseemingly contradictory conclusions with one finding that it results in asensitivity that is a factor of 2 low and the other concluding that it isaccurate. The present work is meant to add to this discussion by exploringthe extent to which <span class="inline-formula">O<sub>3</sub></span> photodissociates in the PAS cell and the rolethat the identity of the bath gas plays in determining the PAS sensitivity.We find a 5<span class="thinspace"></span>% loss in PAS signal attributable to photodissociation at 532<span class="thinspace"></span>nmin <span class="inline-formula">N<sub>2</sub></span> but no loss in a 5<span class="thinspace"></span>% mixture of <span class="inline-formula">O<sub>2</sub></span> in <span class="inline-formula">N<sub>2</sub></span>.Furthermore, we discovered a dramatic increase of more than a factor of 2in the PAS sensitivity as we increased the <span class="inline-formula">O<sub>2</sub></span> fraction in the bathgas, which reached an asymptote near 100<span class="thinspace"></span>% <span class="inline-formula">O<sub>2</sub></span> that nearly matched thesensitivity measured with both <span class="inline-formula">NO<sub>2</sub></span> and nigrosin particles. Weinterpret this dependence with a kinetic model that suggests the reason forthe observed results is a more efficient transfer of energy from excited<span class="inline-formula">O<sub>3</sub></span> to <span class="inline-formula">O<sub>2</sub></span> than to <span class="inline-formula">N<sub>2</sub></span> by a factor of 22–55 depending onexcitation wavelength. Notably, the two prior studies on this topic useddifferent bath gas compositions, and although the results presented here donot fully resolve the differences in their results, they may at leastpartially explain them.</p> | National Science Foundation | 1638307 |  | 2018 |  |  |  | Journal Article